## Specification
Appendix No. 4 - Technical specifications and price calculation
Name of device/equipment/product | Number of pieces | Participant's offer (Manufacturer and type) | Minimum technical specifications, parameters and requirements for the content of the offer as requested by the contracting authority | Parameter compliance | Parameter value for the subject of performance offered by the participant | Price per unit (in EUR without VAT) | Price per unit (in EUR with VAT)
1.- Mobile monitoring station – trailer version | 1 | [seller to add manufacturer, type and year of manufacture] | The supplier shall obtain and deliver all necessary approvals and certifications required for the use of the measuring trailer in normal road traffic, including the vehicle registration certificate and registration plates (state registration plates). For use with driving licence category B, in accordance with Directive 2006/126/EC of the European Parliament. | YES/NO
|  |  | The station must be equipped with air conditioning and heating. | YES/NO
|  |  | A connection cable with a minimum length of 50 m on a reel with all standard connectors used in the Czech Republic for connection to a 230 V and 400 V network. Cable placement inside the monitoring station with secure cable access. | YES/NO | [to be completed by the seller]
|  |  | The measuring trailer must be equipped with an uninterruptible power supply (UPS) with an output voltage of AC 230 V, 50 Hz and a minimum load of 3000 VA for system backup. | YES/NO | [to be completed by the seller]
|  |  | Internal electrical installation: wiring, including the circuit breaker box, must be located inside, 230/400 V switchboard, electricity meter for measuring electricity consumption, sufficient number of socket circuits. Electrical and gas wiring must be located in ducts that allow easy access to these wiring systems. | YES/NO
|  |  | The station must have roof bushings for sampling probes. | YES/NO
|  |  | Maximum trailer weight 750 kg | YES/NO | [to be completed by the seller]
|  |  | The walls of the container and isolation must be made of materials that are inert to the measured parameters. | YES/NO
|  |  | Placement of analyzers and other measuring equipment in 19˝ racks for fixed placement of analyzers. | YES/NO
|  |  | Equipping the monitoring station with basic furniture (chair, table, and shelf for storing measurement documentation) | YES/NO | [to be completed by the seller]
|  |  | Monitoring station cooling (air conditioning capacity) at least 2200 W | YES/NO | [to be completed by the seller]
|  |  | Heating of the monitoring station with a heating element with a minimum output of 500 W | YES/NO | [to be completed by the seller]
2.- Sampling probe for gaseous pollutants | 1 | [to be completed by the seller] | The sampling probe is designed to collect air pollutants and transport them to the connection point of individual analysers for measuring gaseous pollutants. Air supply with a glass manifold with at least 8 outlets, unused routes must be capable of being blinded with a cover, all parts coming into contact with the sample must be made only of borosilicate glass, Teflon (PTFE) or stainless steel, The inlet of the sampling probe must be protected against precipitation and insects entering the measuring system, a fan or pump to ensure a sample flow rate of 2 - 30 l/min, including measurement and regulation of the sample flow. | YES/NO | [to be completed by the seller]
3.- NO, NOx, NO2 analyser based on chemiluminescence | 1 | [to be completed by the seller] | Type approval: Certificate from a reputable European accredited laboratory, the analyser complies with the requirements of the ČSN EN 14211 standard for NO-NO2-NOx (Czech version of European standard EN 14211-2012). | YES/NO
|  |  | Certificate (type approval): To be attached to the offer. | YES/NO
|  |  | Automatic calibration at a frequency corresponding to the parameters of the supplied analyzer. | YES/NO
|  |  | Start-up: Automatic start-up after power failure. | YES/NO
4.- Dust analyser based on optical principle | 1 | [to be completed by the seller] | Simultaneous continuous measurement of particles – TSP, PM10, PM2.5 and PM1 fractions and particle concentration (number of particles) per unit volume. | YES/NO
|  |  | Equivalence test: Certified by a renowned European accredited laboratory, the analyser fully complies with the requirements of standards ČSN EN 16450:2017, ČSN EN 15267-1:2009, ČSN EN 15267-2:2009. | YES/NO
|  |  | Certificate (type approval): To be attached to the offer. | YES/NO
|  |  | Calibration: The calibration process must not interrupt the continuity of measurements, including hourly averages, throughout the entire monitoring period. The device must allow for manual calibration at the measurement site. | YES/NO
5.- Automatic sequential medium-volume sampler | 1 | [to be completed by the seller] | Design: indoor up to 19˝rack, in a stainless steel box. | YES/NO
|  |  | Standard requirements: certification according to ČSN EN 12341. | YES/NO
|  |  | Automatic sampling mode: the sampler must have a filter reservoir (changer) that allows automatic filter replacement without operator intervention. | YES/NO
|  |  | Connection of the device to the data collection and processing system: the supplier shall ensure the connection of the new sampling device to the customer's existing system for the collection and processing of measured data, including the connection and configuration of the station software and the central station. | YES/NO
6.- Calibration system | 1 | [to be completed by the seller] | The system supplier shall ensure the initial filling of the pressure cylinder and supply the appropriate two-stage pressure reducing valve. Min. 10 l cylinder with a mixture of calibration gases NO + SO2 + CO in synthetic air at a concentration of 2/3 of the analyser range, two-stage pressure reducing valve. The supplier shall also ensure the supply of zero gas for daily automatic zero checking. | YES/NO | [to be completed by the seller]
7.- Meteorological components | 1 | [to be completed by the seller] | Meteorological mast: meteorological sensors will be attached to the mast using an extendable arm (length min. 3.5 m when fully extended). | YES/NO | [to be completed by the seller]
|  |  | Wind speed and direction - ultrasonic anemometer. | YES/NO
|  |  | All components shall be corrosion-resistant. | YES/NO
8.- Sensor for measuring global radiation - pyranometer | 1 | [to be completed by the seller] | Calibration certificate included in delivery. | YES/NO
9.- Automatic analyser for measuring ozone (O3), including internal ozone generator based on ultraviolet absorption photometry | 1 | [to be completed by the seller] | Measuring method: Ultraviolet absorption photometry. | YES/NO
|  |  | Type approval: certificate from a renowned European accredited laboratory, the analyser complies with the requirements of the ČSN EN 14625 standard for O3 (Czech version of European standard EN 14625:2012). | YES/NO
|  |  | Automatic calibration at a frequency corresponding to the parameters of the supplied analyzer. | YES/NO
|  |  | Certificate (type approval): To be attached to the offer. | YES/NO
|  |  | Other: Set of consumable parts for 1 year of operation. | YES/NO
10.- SO2 analyser based on UV fluorescence | 1 | [to be completed by the seller] | Measuring method: UV fluorescence. | YES/NO
|  |  | Type approval: certificate from a renowned European accredited laboratory, the analyser complies with the requirements of the ČSN EN 14212 standard for SO2 (Czech version of the European standard EN 14212-2012). | YES/NO
|  |  | Automatic calibration at a frequency corresponding to the parameters of the supplied analyzer. | YES/NO
|  |  | Certificate (type approval): To be attached to the offer. | YES/NO
11.- CO analyser based on non-dispersive infrared spectroscopy | 1 | [to be completed by the seller] | Type approval: certificate from a renowned European accredited laboratory, the analyzer complies with the requirements of the ČSN EN 14626 standard for CO (Czech version of the European standard EN 14626-2012). | YES/NO
|  |  | Automatic calibration at a frequency corresponding to the parameters of the supplied analyzer. | YES/NO
|  |  | Certificate (type approval): To be attached to the offer. | YES/NO
12.- Low-volume sequential sampling device for sampling on sorption tubes | 1 | [to be completed by the seller] | The device fully complies with the requirements of standard ČSN EN 14662-1:2005. | YES/NO
|  |  | Sampling device for taking gas samples from the ambient air for VOC analysis (including benzene). | YES/NO
13.- Datalogger – station control unit | 1 | [to be completed by the seller] | Enables the recording and storage of measured data from all measuring devices – RS232, analogue and LAN inputs. | YES/NO
|  |  | Remote control and automatic and manual calibration start | YES/NO
|  |  | Two-way communication with the central station via a GPRS modem. | YES/NO
14.- Central station | 1 | [to be completed by the seller] | Office PC including accessories. | YES/NO
|  |  | Operating system suitable for the supplied software, including antivirus program. | YES/NO
|  |  | Software for the central station for data management and display. The measured data will be in ISO 7168 format. | YES/NO
|  |  | The measured data should be displayable in both graphical and tabular form, including comparison with emission limits. | YES/NO
|  |  | Minute, minimum 15-minute, half-hour, hourly and daily averages of the measured data will be displayed. In the case of ozone, 8-hour moving averages will also be displayed. The central station will graphically display the measured wind direction values using a wind rose. | YES/NO
15. Project documentation, installation, operator training, laboratory analyses | 1 | [to be completed by the seller] | Installation of all systems in the form of a turnkey delivery. | YES/NO
|  |  | Operator training for at least 4 working days. | YES/NO
|  |  | Ensuring calibration and servicing by the supplier for the duration of the contract. | YES/NO
|  |  | Project documentation for all installed systems. | YES/NO
TOTAL PURCHASE PRICE IN EUR | [5 empty] | 0 | 0
Additional requirements |  |  | Specification of the contracting authority's requirements | Compliance with requirements | Comments
16.- General - common requirements |  |  | The measuring trailer and all its equipment must comply with all health, fire (including powder fire extinguishers) and safety regulations, including sound insulation and the placement of pressure cylinders in holders. | YES/NO
|  |  | Total maximum weight of the measuring trailer, including all supplied equipment and accessories (the seller shall add the total maximum weight of the trailer (measuring trailer including all supplied equipment and accessories) in tonnes) |  | [to be completed by the seller]
|  |  | The following applies to items 3, 4, 9, 10 and 11: - Display: touchscreen; the basic screen shows at least the measured concentration value in the required resolution and alarm indications, with a menu displayed in English with the option of listing all basic parameters that affect the calibration of the meter. - Diagnostics: Comprehensive control of analyser parameters, option of manual setting of measuring device calibration parameters, calculation of minute averages, full internal and remote diagnostics for all essential analyser functions, display of all basic values and diagnostics on the PC screen and remote control of all functions. - Communication: Ethernet/LAN. Option of remote control, remote control of basic functions and option of remote data collection and additional data transfer from the measuring device memory. | YES/NO
|  |  | Documentation: Complete maintenance and operating instructions for the delivered product in printed and electronic form in Czech, including pneumatic and electrical diagrams and a list of basic components and spare parts. | YES/NO
|  |  | Data storage in internal memory: At least monthly continuous measurement. | YES/NO
|  |  | Automatic calibration at a frequency corresponding to the parameters of the supplied analyzer. | YES/NO
|  |  | Warranty and free warranty service for a minimum of 24 months from the date of delivery (acceptance of the purchased item) | YES/NO | [the seller shall specify the warranty period in months]
|  |  | Guaranteed after-sales service and availability of spare parts for at least 5 years after the end of the warranty period | YES/NO